## Slide 1
Lunar calibration of the O2 absorbing channels: 688 and 764 nm
A. Marshak, M. Kowalewski, M. Sutton

## Slide 2
Lunar calibration of the O2 absorbing channels: 688 and 764 nm
Assumptions:
Lunar reflection is constant relative to the adjacent pairs: 680 and 688; 764 and 780;
The ratio between two neighboring channels doesn’t depend on Solar and Viewing angles

## Slide 3
Assumptions:
Lunar reflection is constant relative to the adjacent pairs: 680 and 688; 764 and 780
Using GOME Lunar Albedo:
10 nm difference ~ 0.5% increase
10 nm difference ~ 1.0-1.5% increase
Lunar calibration of the O2 absorbing channels: 688 and 764 nm

## Slide 4
Assumptions:
Lunar reflection is constant relative to the adjacent pairs: 680 and 688; 764 and 780
Lunar calibration of the O2 absorbing channels: 688 and 764 nm
- Moon reflectance Rλ increases slowly with λ; in most cases, a 10 nm difference in wavelength leads to a difference in reflectance in the range of 0.0006-0.0013 or 0.8-1.2%.
- The difference in Rλ between 688 nm and 680 nm as well as between 764 nm and 780 nm will be within 1.6%.
- We use the moon refl. ratios R688/R680 = 1.008 and R764/R780 = 0.984.

## Slide 5
Lunar calibration of the O2 abs. channels;
different solar and viewing angles
DOY 189-191
from Matt Kowalewski & Alexander Cede

## Slide 6
Lunar calibration of the O2 abs. channels
< 1%
Also, the difference between the ROLO ratios and the Lunar derived ratios are ~1.5%

## Slide 7
Lunar calibration of the O2 abs. channels
Examples of moon observations in 680 (left) and 688 (right) on 09/01/2016 (courtesy of Matt Kowalewski).

## Slide 8
Lunar calibration of the O2 abs. channels
Ratios of Moon reflectance F in counts/sec at 688 over 680 nm and at 764 over 780 nm channels.
F(680,688)
and
F(780,764)

## Slide 9: 780 vs. 764
Courtesy of Marshall Sutton

## Slide 10
Lunar calibration of the O2 abs. channels
17 images without the stray light cor. and 36 images with the stray light cor.
The difference in 764/780 ratio < 2%, and the difference in the 688/680 ratio < 1%.
F(680,688)
and
F(780,764)
Courtesy of Marshall Sutton
new moon

## Slide 11
Lunar calibration of the O2 abs. channels
K688 = R688/R688counts
= R688/[R680counts F(680,688)]
= [R688 K680]/[R680 F(680,688)]
= [R688/R680] [K680/F(680,688)]
≈ 1.008 [K680/F(680,688)] = 2.02 10-5
K764 ≈ 0.984 [K780/F(780,764)] = 2.36 10-5

## Slide 12
To covert Level 1 data into reflectance for each λ we have to multiply data given in engineering units (counts/second) by calibration factors Kλ. The accuracy of the calibration factors is about 1-3%.

## Slide 13
Thank you

## Slide 15: R05-V02: 09/01/2016 - 680nm
Examples of moon observations in 680 (left) and 688 (right) on 09/01/2016 (courtesy of Matthew Kowalewski).

## Slide 16: EPIC_1a_20160719143037_01.h5 780nm / 764nm
Square Size | OLS Forced Through Origin |  | OLS |  |  | Average
 | Slope | Variance | Slope | Intercept | Variance | Average
2 | 0.5962 | 0.0005 | 0.5613 | 304.29 | 0.002 | 0.6005
4 | 0.5969 | 0.0008 | 0.5731 | 207.48 | 0.003 | 0.5998
16 | 0.5972 | 0.0018 | 0.5932 | 34.35 | 0.009 | 0.5976
32 | 0.5976 | 0.0024 | 0.6040 | -54.911 | 0.013 | 0.5972
64 | 0.5981 | 0.0031 | 0.6286. | -257.43 | 0.017 | 0.5966